National Weights & Measures Lab, Delhi

Sector 5, Metrology Bhawan, New Delhi

Report No.: SS-2026-0007    Date: 2026-09-26 14:32

# Type Evaluation / Verification Report

## Instrument Under Test

| Manufacturer | Accura Scales Pvt. Ltd. |
| --- | --- |
| Model | AX-500 |
| Serial No. | AX500-2291 |
| Accuracy Class | III |
| Max Capacity | 5000 g |
| Min Capacity | 20 g |
| Scale Interval (e) | 1 g |
| Verification Scale Interval (d) | 1 g |
| Max Tare | 2000 g |

## Test Conditions

| Test Location | Lab Bay 2 |
| --- | --- |
| Tested On | 2026-09-25 11:00 |
| Tested By | Rahul Mehta |
| Temperature | 24.5 °C |
| Humidity | 55 %RH |
| Pressure | 1004 hPa |

⚠ Rounding-trap detected. The naive calculation and the R 76-1 changeover-point method disagree on the verdict for at least one reading below. The R 76-1 result is authoritative.

Marginal result(s). The following readings used more than 80% of the allowed MPE.

## Test Results

| Test | Load (g) | Error (g) | MPE (±g) | MPE Used | Naive rounding | Result | Clause | Explanation |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| weighing | 1000 | -3.0 | 2.5 | 120% | 2.0 → PASS | FAIL | R76-1 3.6.1 | Error of -3.0 g exceeds MPE of ±2.5 g for loads between 500e and 2000e. |
| eccentricity | 2500 | 1.0 | 5.0 | 20% | 1.0 → PASS | PASS | R76-1 3.6.2 | Corner-load error within MPE. |
| repeatability ⚠ | 5000 | 2.1 | 2.5 | 84% | 2.1 → PASS | PASS | R76-1 3.6.3 | Spread of repeat readings uses 84% of MPE (marginal). |

Overall Verdict: FAIL

| Dr. Anita Rao Chief Metrology Officer Approved On: 2026-09-26 14:32 |  |
| --- | --- |

Scan the QR code or visit the link below to verify this report is genuine and unmodified. https://scalesaathi.example.org/verify/3f9a7c1e2b8d4f0a6c5e9b1d7a2f4c8e0b6d3a1f9c7e5b2d4a6f8c0e2b4d6f81
This report is system-generated and cryptographically hashed at approval. Any change after approval will be flagged at the verification link above.